## Summary
DEV-00196, ATTACHMENT X, Version Y | [11 empty] | SINGLE CUBE CALIBRABRATION DATA EVALUATION
| Well # | CUBE HOST NAME==> |  |  | B2030000A | R2 Calculations | [5 empty] | B19460006
Cartridge Ref |  | Key Results | Target |  | Individual Reading CUBE #1 | Dil. | Decimal Dilution | Well Average Fl. | R2 Target |  | R2 Value | Optical Standards Readings From REFERENCE CUBE
|  |  | Min | Max |  |  |  |  | Min | Max
16X Dilution Cartridge 1 GREEN & Cartridge 4 RED | Well #1 | Post CAL Green Photo 1 | 20564 | 37820 | 31576.5915 | 16x | 0.0625 | 31607.6521 | 0.994 | 1 | 0.99802 | 31552.1055
|  | Post CAL Green Photo 2 | 20564 | 37820 | 31616.2598 | 32x | 0.03125 | 16150.9074
|  | Post CAL Green Photo 3 | 20564 | 37820 | 31630.1049 | 64x | 0.015625 | 6955.8732 |  |  | Well #1 GREEN R2
| Well #2 | Post CAL Green Photo 1 | 22076 | 36890 | 32658.5511 | 16x | 0.0625 | 32673.8804 | 0.994 | 1 | 0.99952 | 32675.1877
|  | Post CAL Green Photo 2 | 22076 | 36890 | 32681.6161 | 32x | 0.03125 | 16510.0302
|  | Post CAL Green Photo 3 | 22076 | 36890 | 32681.474 | 64x | 0.015625 | 7691.7596 |  |  | Well #2 GREEN R2
| Well #3 | Post CAL Green Photo 1 | 21750 | 35256 | 30237.568 | 16x | 0.0625 | 30269.5039 | 0.994 | 1 | 0.99978 | 30245.7547
|  | Post CAL Green Photo 2 | 21750 | 35256 | 30288.2878 | 32x | 0.03125 | 15051.5815
|  | Post CAL Green Photo 3 | 21750 | 35256 | 30282.656 | 64x | 0.015625 | 6976.8482 |  |  | Well #3 GREEN R2
| Well #1 | Post CAL RED Photo 1 | 24525 | 33579 | 30396.6914 | 16x | 0.0625 | 30395.0354 | 0.994 | 1 | 0.99959 | 30432.9104
|  | Post CAL RED Photo 2 | 24525 | 33579 | 30395.9644 | 32x | 0.03125 | 14132.7306
|  | Post CAL RED Photo 3 | 24525 | 33579 | 30392.4504 | 64x | 0.015625 | 6649.8324 |  |  | Well #1 RED R2
| Well #2 | Post CAL RED Photo 1 | 23185 | 34885 | 33805.284 | 16x | 0.0625 | 33801.8661 | 0.994 | 1 | 0.99971 | 33696.0392
|  | Post CAL RED Photo 2 | 23185 | 34885 | 33816.328 | 32x | 0.03125 | 16017.6665
|  | Post CAL RED Photo 3 | 23185 | 34885 | 33783.9863 | 64x | 0.015625 | 6496.3867 |  |  | Well #2 RED R2
| Well #3 | Post CAL RED Photo 1 | 25189 | 32761 | 30135.6806 | 16x | 0.0625 | 30131.2888 | 0.994 | 1 | 0.99988 | 30077.4416
|  | Post CAL RED Photo 2 | 25189 | 32761 | 30126.7992 | 32x | 0.03125 | 13988.9438
|  | Post CAL RED Photo 3 | 25189 | 32761 | 30131.3865 | 64x | 0.015625 | 6277.6575 |  |  | Well #3 RED R2
| Well #1 (Cart. #1) | Post CAL BLUE Photo 1 | 22000 | 33000 | 242927.5556
|  | Post CAL BLUE Photo 2 | 22000 | 33000 | 243441.5556
|  | Post CAL BLUE Photo 3 | 22000 | 33000 | 243415.0556
| Well #2 (Cart. #1) | Post CAL BLUE Photo 1 | 22000 | 33000 | 277211.6667
|  | Post CAL BLUE Photo 2 | 22000 | 33000 | 277549.6667
|  | Post CAL BLUE Photo 3 | 22000 | 33000 | 277813.1667
| Well #3 (Cart. #1) | Post CAL BLUE Photo 1 | 22000 | 33000 | 238783.2222
|  | Post CAL BLUE Photo 2 | 22000 | 33000 | 238820.2222
|  | Post CAL BLUE Photo 3 | 22000 | 33000 | 238822.7222
| Well #1 (Cart. #4) | Post CAL BLUE Photo 1 | 22000 | 33000 | 39230.0556
|  | Post CAL BLUE Photo 2 | 22000 | 33000 | 39136.5556
|  | Post CAL BLUE Photo 3 | 22000 | 33000 | 38936.0556
| Well #2 (Cart. #4) | Post CAL BLUE Photo 1 | 22000 | 33000 | 39680.1667
|  | Post CAL BLUE Photo 2 | 22000 | 33000 | 39699.6667
|  | Post CAL BLUE Photo 3 | 22000 | 33000 | 39773.1667
| Well #3 (Cart. #4) | Post CAL BLUE Photo 1 | 22000 | 33000 | 35066.7222
|  | Post CAL BLUE Photo 2 | 22000 | 33000 | 35043.2222
|  | Post CAL BLUE Photo 3 | 22000 | 33000 | 34956.7222
32X Dilution Cartridge 2 GREEN & Cartridge 5 RED | Well #1 | Post CAL Green Photo 1 | 9578 | 18860 | 16136.7796 | [6 empty] | 16213.0861
|  | Post CAL Green Photo 2 | 9578 | 18860 | 16150.9608
|  | Post CAL Green Photo 3 | 9578 | 18860 | 16164.9817
| Well #2 | Post CAL Green Photo 1 | 10489 | 18667 | 16504.2843 | [6 empty] | 16322.6945
|  | Post CAL Green Photo 2 | 10489 | 18667 | 16511.0629
|  | Post CAL Green Photo 3 | 10489 | 18667 | 16514.7433
| Well #3 | Post CAL Green Photo 1 | 9980 | 18134 | 15041.7743 | [6 empty] | 15356.9713
|  | Post CAL Green Photo 2 | 9980 | 18134 | 15051.0792
|  | Post CAL Green Photo 3 | 9980 | 18134 | 15061.8909
| Well #1 | Post CAL RED Photo 1 | 9762 | 17118 | 14129.5469 | [6 empty] | 14191.0297
|  | Post CAL RED Photo 2 | 9762 | 17118 | 14141.8121
|  | Post CAL RED Photo 3 | 9762 | 17118 | 14126.8329
| Well #2 | Post CAL RED Photo 1 | 10224 | 17406 | 16015.1264 | [6 empty] | 16213.1043
|  | Post CAL RED Photo 2 | 10224 | 17406 | 16026.0289
|  | Post CAL RED Photo 3 | 10224 | 17406 | 16011.8442
| Well #3 | Post CAL RED Photo 1 | 9865 | 17017 | 13988.94 | [6 empty] | 13806.2851
|  | Post CAL RED Photo 2 | 9865 | 17017 | 14001.4313
|  | Post CAL RED Photo 3 | 9865 | 17017 | 13976.46
| Well #1 (Cart. #2) | Post CAL BLUE Photo 1 | 11000 | 16500 | 122772.0556
|  | Post CAL BLUE Photo 2 | 11000 | 16500 | 122855.0556
|  | Post CAL BLUE Photo 3 | 11000 | 16500 | 123064.0556
| Well #2 (Cart. #2) | Post CAL BLUE Photo 1 | 11000 | 16500 | 138384.1667
|  | Post CAL BLUE Photo 2 | 11000 | 16500 | 138432.1667
|  | Post CAL BLUE Photo 3 | 11000 | 16500 | 138538.1667
| Well #3 (Cart. #2) | Post CAL BLUE Photo 1 | 11000 | 16500 | 117637.7222
|  | Post CAL BLUE Photo 2 | 11000 | 16500 | 117481.7222
|  | Post CAL BLUE Photo 3 | 11000 | 16500 | 117617.7222
| Well #1 (Cart. #5) | Post CAL BLUE Photo 1 | 11000 | 16500 | 18430.5556
|  | Post CAL BLUE Photo 2 | 11000 | 16500 | 18412.5556
|  | Post CAL BLUE Photo 3 | 11000 | 16500 | 18387.5556
| Well #2 (Cart. #5) | Post CAL BLUE Photo 1 | 11000 | 16500 | 18993.6667
|  | Post CAL BLUE Photo 2 | 11000 | 16500 | 18846.6667
|  | Post CAL BLUE Photo 3 | 11000 | 16500 | 18708.6667
| Well #3 (Cart. #5) | Post CAL BLUE Photo 1 | 11000 | 16500 | 16411.2222
|  | Post CAL BLUE Photo 2 | 11000 | 16500 | 16401.2222
|  | Post CAL BLUE Photo 3 | 11000 | 16500 | 16307.2222
64X Dilution Cartridge 3 GREEN & Cartridge 6 RED | Well #1 | Post CAL Green Photo 1 | 3752 | 9080 | 6944.4105 | [6 empty] | 6727.4892
|  | Post CAL Green Photo 2 | 3752 | 9080 | 6963.9037
|  | Post CAL Green Photo 3 | 3752 | 9080 | 6959.3053
| Well #2 | Post CAL Green Photo 1 | 3909 | 9105 | 7689.1675 | [6 empty] | 7739.0747
|  | Post CAL Green Photo 2 | 3909 | 9105 | 7693.7058
|  | Post CAL Green Photo 3 | 3909 | 9105 | 7692.4054
| Well #3 | Post CAL Green Photo 1 | 3543 | 8811 | 6965.1258 | [6 empty] | 7202.2284
|  | Post CAL Green Photo 2 | 3543 | 8811 | 6985.6622
|  | Post CAL Green Photo 3 | 3543 | 8811 | 6979.7567
| Well #1 | Post CAL RED Photo 1 | 3505 | 7885 | 6656.216 | [6 empty] | 6598.9444
|  | Post CAL RED Photo 2 | 3505 | 7885 | 6646.3781
|  | Post CAL RED Photo 3 | 3505 | 7885 | 6646.9031
| Well #2 | Post CAL RED Photo 1 | 3619 | 8053 | 6491.2057 | [6 empty] | 6094.7451
|  | Post CAL RED Photo 2 | 3619 | 8053 | 6496.0189
|  | Post CAL RED Photo 3 | 3619 | 8053 | 6501.9354
| Well #3 | Post CAL RED Photo 1 | 3365 | 7853 | 6280.4589 | [6 empty] | 6107.8198
|  | Post CAL RED Photo 2 | 3365 | 7853 | 6283.1127
|  | Post CAL RED Photo 3 | 3365 | 7853 | 6269.4009
| Well #1 (Cart. #3) | Post CAL BLUE Photo 1 | 5500 | 8250 | 53923.0556
|  | Post CAL BLUE Photo 2 | 5500 | 8250 | 53949.0556
|  | Post CAL BLUE Photo 3 | 5500 | 8250 | 54030.0556
| Well #2 (Cart. #3) | Post CAL BLUE Photo 1 | 5500 | 8250 | 64817.1667
|  | Post CAL BLUE Photo 2 | 5500 | 8250 | 64905.1667
|  | Post CAL BLUE Photo 3 | 5500 | 8250 | 64806.1667
| Well #3 (Cart. #3) | Post CAL BLUE Photo 1 | 5500 | 8250 | 55228.7222
|  | Post CAL BLUE Photo 2 | 5500 | 8250 | 55254.7222
|  | Post CAL BLUE Photo 3 | 5500 | 8250 | 55307.7222
| Well #1 (Cart. #6) | Post CAL BLUE Photo 1 | 5500 | 8250 | 9429.5556
|  | Post CAL BLUE Photo 2 | 5500 | 8250 | 9578.5556
|  | Post CAL BLUE Photo 3 | 5500 | 8250 | 9470.0556
| Well #2 (Cart. #6) | Post CAL BLUE Photo 1 | 5500 | 8250 | 8606.6667
|  | Post CAL BLUE Photo 2 | 5500 | 8250 | 8566.6667
|  | Post CAL BLUE Photo 3 | 5500 | 8250 | 8632.1667
| Well #3 (Cart. #6) | Post CAL BLUE Photo 1 | 5500 | 8250 | 8201.2222
|  | Post CAL BLUE Photo 2 | 5500 | 8250 | 8146.2222
|  | Post CAL BLUE Photo 3 | 5500 | 8250 | 8355.7222
Product Engineering Review and Approval
Quality Assurance Review and Approval
Page 1 of 1

## Buffers
DEV-00196, ATTACHMENT X, Version Y | [5 empty] | SINGLE CUBE CALIBRABRATION DATA EVALUATION
| Well # | CUBE HOST NAME==> |  |  | B2030000A
Cartridge Ref |  | Key Results | Target |  | Individual Reading CUBE #1
|  |  | Min | Max
Cartridge 7 Buffer | Well #1 | Post CAL BUFFER Photo 1 | -500 | 500 | 4.29507201011951
|  | Post CAL BUFFER Photo 2 | -500 | 500 | 8.82156687743035
|  | Post CAL BUFFER Photo 3 | -500 | 500 | 10.0184802192698
| Well #2 | Post CAL BUFFER Photo 1 | -500 | 500 | -10.1114627286696
|  | Post CAL BUFFER Photo 2 | -500 | 500 | -3.48288922911448
|  | Post CAL BUFFER Photo 3 | -500 | 500 | -6.43285913620193
| Well #3 | Post CAL BUFFER Photo 1 | -500 | 500 | 7.93276652418206
|  | Post CAL BUFFER Photo 2 | -500 | 500 | 12.111466600377
|  | Post CAL BUFFER Photo 3 | -500 | 500 | 9.76099582012333
| Well #1 | Post CAL BUFFER Photo 1 | -500 | 500 | -16.7071142246698
|  | Post CAL BUFFER Photo 2 | -500 | 500 | -14.8483627479272
|  | Post CAL BUFFER Photo 3 | -500 | 500 | -17.2250973021501
| Well #2 | Post CAL BUFFER Photo 1 | -500 | 500 | -8.28730160416204
|  | Post CAL BUFFER Photo 2 | -500 | 500 | -7.27607232956695
|  | Post CAL BUFFER Photo 3 | -500 | 500 | -4.79627084421514
| Well #3 | Post CAL BUFFER Photo 1 | -500 | 500 | -8.89285131595874
|  | Post CAL BUFFER Photo 2 | -500 | 500 | -4.30105102022712
|  | Post CAL BUFFER Photo 3 | -500 | 500 | -6.71949522973905
Cartridge 8 Buffer | Well #1 | Post CAL BUFFER Photo 1 | -500 | 500 | 44.6902417262101
|  | Post CAL BUFFER Photo 2 | -500 | 500 | -32.0433563270785
|  | Post CAL BUFFER Photo 3 | -500 | 500 | -23.1423623328577
| Well #2 | Post CAL BUFFER Photo 1 | -500 | 500 | 21.0248757976513
|  | Post CAL BUFFER Photo 2 | -500 | 500 | -61.7596227360597
|  | Post CAL BUFFER Photo 3 | -500 | 500 | -55.5204417953464
| Well #3 | Post CAL BUFFER Photo 1 | -500 | 500 | 3.07591764414768
|  | Post CAL BUFFER Photo 2 | -500 | 500 | -79.9979541588013
|  | Post CAL BUFFER Photo 3 | -500 | 500 | -75.2525006813337
| Well #1 | Post CAL BUFFER Photo 1 | -500 | 500 | 22.9855583372967
|  | Post CAL BUFFER Photo 2 | -500 | 500 | 5.95457516386857
|  | Post CAL BUFFER Photo 3 | -500 | 500 | 13.213106818688
| Well #2 | Post CAL BUFFER Photo 1 | -500 | 500 | 8.0134856653508
|  | Post CAL BUFFER Photo 2 | -500 | 500 | -9.19769823845091
|  | Post CAL BUFFER Photo 3 | -500 | 500 | -7.46650267502834
| Well #3 | Post CAL BUFFER Photo 1 | -500 | 500 | 5.97132450590768
|  | Post CAL BUFFER Photo 2 | -500 | 500 | -5.94718703455439
|  | Post CAL BUFFER Photo 3 | -500 | 500 | -7.45888527438266
Cartridge 9 Buffer | Well #1 | Post CAL BUFFER Photo 1 | -500 | 500 | -44.0364906144248
|  | Post CAL BUFFER Photo 2 | -500 | 500 | 57.3167398551136
|  | Post CAL BUFFER Photo 3 | -500 | 500 | -25.9198914137822
| Well #2 | Post CAL BUFFER Photo 1 | -500 | 500 | -4.10902507477282
|  | Post CAL BUFFER Photo 2 | -500 | 500 | 101.724905478558
|  | Post CAL BUFFER Photo 3 | -500 | 500 | 18.666519423956
| Well #3 | Post CAL BUFFER Photo 1 | -500 | 500 | 1.13135119744173
|  | Post CAL BUFFER Photo 2 | -500 | 500 | 96.3311899869857
|  | Post CAL BUFFER Photo 3 | -500 | 500 | 24.9067670668778
| Well #1 | Post CAL BUFFER Photo 1 | -500 | 500 | 24.112754671342
|  | Post CAL BUFFER Photo 2 | -500 | 500 | -13.3092450307394
|  | Post CAL BUFFER Photo 3 | -500 | 500 | -4.17617568570892
| Well #2 | Post CAL BUFFER Photo 1 | -500 | 500 | 30.6577157325565
|  | Post CAL BUFFER Photo 2 | -500 | 500 | -4.55073563982736
|  | Post CAL BUFFER Photo 3 | -500 | 500 | 2.90337993334345
| Well #3 | Post CAL BUFFER Photo 1 | -500 | 500 | 30.7964688782677
|  | Post CAL BUFFER Photo 2 | -500 | 500 | -6.46747916444341
|  | Post CAL BUFFER Photo 3 | -500 | 500 | 3.01915565512972
Product Engineering Review and Approval
Quality Assurance Review and Approval

## #1 Cube Results
Insert Cube Hostname Here==> | B2030000A
|  | PRE CALIBRATION | STANDARD #1
|  | PHOTO #1 | GREEN 16X
146962 |  | RED WELL #1 | 146962
169754 |  | RED WELL #2 | 169754
141195 |  | RED WELL #3 | 141195
766650.5 |  | GREEN WELL #1 | 766650.5
939531.5 |  | GREEN WELL #2 | 939531.5
768666.5 |  | GREEN WELL #3 | 768666.5
254859.7778 |  | BLUE WELL #1 | 254859.777777778
290441.7778 |  | BLUE WELL #2 | 290441.777777778
251215.7778 |  | BLUE WELL #3 | 251215.777777778
STANDARD #1
PHOTO #2
147634 |  | RED WELL #1 | 147634
170300 |  | RED WELL #2 | 170300
141951 |  | RED WELL #3 | 141951
767902.5 |  | GREEN WELL #1 | 767902.5
940484.5 |  | GREEN WELL #2 | 940484.5
770237.5 |  | GREEN WELL #3 | 770237.5
255529.7778 |  | BLUE WELL #1 | 255529.777777778
290935.7778 |  | BLUE WELL #2 | 290935.777777778
251408.7778 |  | BLUE WELL #3 | 251408.777777778
STANDARD #1
PHOTO #3
147542 |  | RED WELL #1 | 147542
170169 |  | RED WELL #2 | 170169
141942 |  | RED WELL #3 | 141942
768204.5 |  | GREEN WELL #1 | 768204.5
940463.5 |  | GREEN WELL #2 | 940463.5
770108.5 |  | GREEN WELL #3 | 770108.5
255767.7778 |  | BLUE WELL #1 | 255767.777777778
291463.7778 |  | BLUE WELL #2 | 291463.777777778
251675.7778 |  | BLUE WELL #3 | 251675.777777778
|  | PRE CALIBRATION | STANDARD #2
|  | PHOTO #1 | GREEN 32X
77758 |  | RED WELL #1 | 77758
88130 |  | RED WELL #2 | 88130
73036 |  | RED WELL #3 | 73036
397115.5 |  | GREEN WELL #1 | 397115.5
481130.5 |  | GREEN WELL #2 | 481130.5
388330.5 |  | GREEN WELL #3 | 388330.5
131288.7778 |  | BLUE WELL #1 | 131288.777777778
148198.7778 |  | BLUE WELL #2 | 148198.777777778
126654.7778 |  | BLUE WELL #3 | 126654.777777778
STANDARD #2
PHOTO #2
77743 |  | RED WELL #1 | 77743
88260 |  | RED WELL #2 | 88260
73225 |  | RED WELL #3 | 73225
397608.5 |  | GREEN WELL #1 | 397608.5
481511.5 |  | GREEN WELL #2 | 481511.5
388759.5 |  | GREEN WELL #3 | 388759.5
131526.7778 |  | BLUE WELL #1 | 131526.777777778
148401.7778 |  | BLUE WELL #2 | 148401.777777778
126653.7778 |  | BLUE WELL #3 | 126653.777777778
STANDARD #2
PHOTO #3
78073 |  | RED WELL #1 | 78073
88507 |  | RED WELL #2 | 88507
73137 |  | RED WELL #3 | 73137
398044.5 |  | GREEN WELL #1 | 398044.5
481714.5 |  | GREEN WELL #2 | 481714.5
389060.5 |  | GREEN WELL #3 | 389060.5
131864.7778 |  | BLUE WELL #1 | 131864.777777778
148636.7778 |  | BLUE WELL #2 | 148636.777777778
126918.7778 |  | BLUE WELL #3 | 126918.777777778
|  | PRE CALIBRATION | STANDARD #3
|  | PHOTO #1 | GREEN 64X
36477 |  | RED WELL #1 | 36477
44481 |  | RED WELL #2 | 44481
36899 |  | RED WELL #3 | 36899
176285.5 |  | GREEN WELL #1 | 176285.5
229967.5 |  | GREEN WELL #2 | 229967.5
184908.5 |  | GREEN WELL #3 | 184908.5
59898.7778 |  | BLUE WELL #1 | 59898.7777777778
72090.7778 |  | BLUE WELL #2 | 72090.7777777778
61704.7778 |  | BLUE WELL #3 | 61704.7777777778
STANDARD #3
PHOTO #2
36690 |  | RED WELL #1 | 36690
44731 |  | RED WELL #2 | 44731
37195 |  | RED WELL #3 | 37195
176840.5 |  | GREEN WELL #1 | 176840.5
230210.5 |  | GREEN WELL #2 | 230210.5
185524.5 |  | GREEN WELL #3 | 185524.5
60226.7778 |  | BLUE WELL #1 | 60226.7777777778
72480.7778 |  | BLUE WELL #2 | 72480.7777777778
62032.7778 |  | BLUE WELL #3 | 62032.7777777778
STANDARD #3
PHOTO #3
36933 |  | RED WELL #1 | 36933
44851 |  | RED WELL #2 | 44851
37293 |  | RED WELL #3 | 37293
176868.5 |  | GREEN WELL #1 | 176868.5
230284.5 |  | GREEN WELL #2 | 230284.5
185489.5 |  | GREEN WELL #3 | 185489.5
60420.7778 |  | BLUE WELL #1 | 60420.7777777778
72494.7778 |  | BLUE WELL #2 | 72494.7777777778
62198.7778 |  | BLUE WELL #3 | 62198.7777777778
|  | PRE CALIBRATION | STANDARD #4
|  | PHOTO #1 | RED 16X
527259 |  | RED WELL #1 | 527259
545618 |  | RED WELL #2 | 545618
508473 |  | RED WELL #3 | 508473
106254.5 |  | GREEN WELL #1 | 106254.5
107835.5 |  | GREEN WELL #2 | 107835.5
99182.5 |  | GREEN WELL #3 | 99182.5
45034.7778 |  | BLUE WELL #1 | 45034.7777777778
46782.7778 |  | BLUE WELL #2 | 46782.7777777778
41371.7778 |  | BLUE WELL #3 | 41371.7777777778
STANDARD #4
PHOTO #2
527531 |  | RED WELL #1 | 527531
546121 |  | RED WELL #2 | 546121
508669 |  | RED WELL #3 | 508669
106331.5 |  | GREEN WELL #1 | 106331.5
108200.5 |  | GREEN WELL #2 | 108200.5
99532.5 |  | GREEN WELL #3 | 99532.5
45206.7778 |  | BLUE WELL #1 | 45206.7777777778
47067.7778 |  | BLUE WELL #2 | 47067.7777777778
41613.7778 |  | BLUE WELL #3 | 41613.7777777778
STANDARD #4
PHOTO #3
527501 |  | RED WELL #1 | 527501
545660 |  | RED WELL #2 | 545660
508747 |  | RED WELL #3 | 508747
106472.5 |  | GREEN WELL #1 | 106472.5
108332.5 |  | GREEN WELL #2 | 108332.5
99565.5 |  | GREEN WELL #3 | 99565.5
45111.7778 |  | BLUE WELL #1 | 45111.7777777778
47246.7778 |  | BLUE WELL #2 | 47246.7777777778
41632.7778 |  | BLUE WELL #3 | 41632.7777777778
|  | PRE CALIBRATION | STANDARD #5
|  | PHOTO #1 | RED 32X
248184 |  | RED WELL #1 | 248184
262108 |  | RED WELL #2 | 262108
239467 |  | RED WELL #3 | 239467
54060.5 |  | GREEN WELL #1 | 54060.5
56684.5 |  | GREEN WELL #2 | 56684.5
50536.5 |  | GREEN WELL #3 | 50536.5
23795.7778 |  | BLUE WELL #1 | 23795.7777777778
25656.7778 |  | BLUE WELL #2 | 25656.7777777778
22276.7778 |  | BLUE WELL #3 | 22276.7777777778
STANDARD #5
PHOTO #2
248569 |  | RED WELL #1 | 248569
262454 |  | RED WELL #2 | 262454
239870 |  | RED WELL #3 | 239870
54189.5 |  | GREEN WELL #1 | 54189.5
56789.5 |  | GREEN WELL #2 | 56789.5
50781.5 |  | GREEN WELL #3 | 50781.5
23942.7778 |  | BLUE WELL #1 | 23942.7777777778
25674.7778 |  | BLUE WELL #2 | 25674.7777777778
22431.7778 |  | BLUE WELL #3 | 22431.7777777778
STANDARD #5
PHOTO #3
245855 |  | RED WELL #1 | 245855
259779 |  | RED WELL #2 | 259779
236984 |  | RED WELL #3 | 236984
54353.5 |  | GREEN WELL #1 | 54353.5
57018.5 |  | GREEN WELL #2 | 57018.5
50822.5 |  | GREEN WELL #3 | 50822.5
23924.7778 |  | BLUE WELL #1 | 23924.7777777778
25543.7778 |  | BLUE WELL #2 | 25543.7777777778
22344.7778 |  | BLUE WELL #3 | 22344.7777777778
|  | PRE CALIBRATION | STANDARD #6
|  | PHOTO #1 | RED 64X
119276 |  | RED WELL #1 | 119276
110388 |  | RED WELL #2 | 110388
110693 |  | RED WELL #3 | 110693
29980.5 |  | GREEN WELL #1 | 29980.5
30042.5 |  | GREEN WELL #2 | 30042.5
28343.5 |  | GREEN WELL #3 | 28343.5
14590.7778 |  | BLUE WELL #1 | 14590.7777777778
15065.7778 |  | BLUE WELL #2 | 15065.7777777778
13862.7778 |  | BLUE WELL #3 | 13862.7777777778
STANDARD #6
PHOTO #2
119416 |  | RED WELL #1 | 119416
110750 |  | RED WELL #2 | 110750
111047 |  | RED WELL #3 | 111047
30201.5 |  | GREEN WELL #1 | 30201.5
30224.5 |  | GREEN WELL #2 | 30224.5
28651.5 |  | GREEN WELL #3 | 28651.5
14897.7778 |  | BLUE WELL #1 | 14897.7777777778
15183.7778 |  | BLUE WELL #2 | 15183.7777777778
13965.7778 |  | BLUE WELL #3 | 13965.7777777778
STANDARD #6
PHOTO #3
119582 |  | RED WELL #1 | 119582
110993 |  | RED WELL #2 | 110993
110979 |  | RED WELL #3 | 110979
30319.5 |  | GREEN WELL #1 | 30319.5
30330.5 |  | GREEN WELL #2 | 30330.5
28711.5 |  | GREEN WELL #3 | 28711.5
14908.7778 |  | BLUE WELL #1 | 14908.7777777778
15368.7778 |  | BLUE WELL #2 | 15368.7777777778
14294.7778 |  | BLUE WELL #3 | 14294.7777777778
|  | PRE CALIBRATION | STANDARD #7
|  | PHOTO #1 | BUFFER #1
5504 |  | RED WELL #1 | 5504
6688 |  | RED WELL #2 | 6688
6376 |  | RED WELL #3 | 6376
9862.5 |  | GREEN WELL #1 | 9862.5
12077.5 |  | GREEN WELL #2 | 12077.5
10699.5 |  | GREEN WELL #3 | 10699.5
5480.7778 |  | BLUE WELL #1 | 5480.77777777778
6778.7778 |  | BLUE WELL #2 | 6778.77777777778
6074.7778 |  | BLUE WELL #3 | 6074.77777777778
STANDARD #7
PHOTO #2
5847 |  | RED WELL #1 | 5847
7055 |  | RED WELL #2 | 7055
6723 |  | RED WELL #3 | 6723
10006.5 |  | GREEN WELL #1 | 10006.5
12210.5 |  | GREEN WELL #2 | 12210.5
10910.5 |  | GREEN WELL #3 | 10910.5
5721.7778 |  | BLUE WELL #1 | 5721.77777777778
7186.7778 |  | BLUE WELL #2 | 7186.77777777778
6477.7778 |  | BLUE WELL #3 | 6477.77777777778
STANDARD #7
PHOTO #3
5932 |  | RED WELL #1 | 5932
7097 |  | RED WELL #2 | 7097
6750 |  | RED WELL #3 | 6750
9977.5 |  | GREEN WELL #1 | 9977.5
12294.5 |  | GREEN WELL #2 | 12294.5
10867.5 |  | GREEN WELL #3 | 10867.5
5849.7778 |  | BLUE WELL #1 | 5849.77777777778
7284.7778 |  | BLUE WELL #2 | 7284.77777777778
6460.7778 |  | BLUE WELL #3 | 6460.77777777778
|  | PRE CALIBRATION | STANDARD #8
|  | PHOTO #1 | BUFFER #2
6492 |  | RED WELL #1 | 6492
7391 |  | RED WELL #2 | 7391
6504 |  | RED WELL #3 | 6504
11014.5 |  | GREEN WELL #1 | 11014.5
12715.5 |  | GREEN WELL #2 | 12715.5
11147.5 |  | GREEN WELL #3 | 11147.5
6254.7778 |  | BLUE WELL #1 | 6254.77777777778
7303.7778 |  | BLUE WELL #2 | 7303.77777777778
6485.7778 |  | BLUE WELL #3 | 6485.77777777778
STANDARD #8
PHOTO #2
3660 |  | RED WELL #1 | 3660
4546 |  | RED WELL #2 | 4546
3631 |  | RED WELL #3 | 3631
9874.5 |  | GREEN WELL #1 | 9874.5
11502.5 |  | GREEN WELL #2 | 11502.5
10113.5 |  | GREEN WELL #3 | 10113.5
3302.7778 |  | BLUE WELL #1 | 3302.77777777778
4312.7778 |  | BLUE WELL #2 | 4312.77777777778
3516.7778 |  | BLUE WELL #3 | 3516.77777777778
STANDARD #8
PHOTO #3
3903 |  | RED WELL #1 | 3903
4714 |  | RED WELL #2 | 4714
3764 |  | RED WELL #3 | 3764
10142.5 |  | GREEN WELL #1 | 10142.5
11638.5 |  | GREEN WELL #2 | 11638.5
10160.5 |  | GREEN WELL #3 | 10160.5
3452.7778 |  | BLUE WELL #1 | 3452.77777777778
4594.7778 |  | BLUE WELL #2 | 4594.77777777778
3751.7778 |  | BLUE WELL #3 | 3751.77777777778
|  | PRE CALIBRATION | STANDARD #9
|  | PHOTO #1 | BUFFER #3
3342 |  | RED WELL #1 | 3342
5435 |  | RED WELL #2 | 5435
4903 |  | RED WELL #3 | 4903
10722.5 |  | GREEN WELL #1 | 10722.5
13232.5 |  | GREEN WELL #2 | 13232.5
11705.5 |  | GREEN WELL #3 | 11705.5
6066.7778 |  | BLUE WELL #1 | 6066.77777777778
7469.7778 |  | BLUE WELL #2 | 7469.77777777778
6582.7778 |  | BLUE WELL #3 | 6582.77777777778
STANDARD #9
PHOTO #2
6973 |  | RED WELL #1 | 6973
9008 |  | RED WELL #2 | 9008
8385 |  | RED WELL #3 | 8385
9684.5 |  | GREEN WELL #1 | 9684.5
12069.5 |  | GREEN WELL #2 | 12069.5
10590.5 |  | GREEN WELL #3 | 10590.5
3172.7778 |  | BLUE WELL #1 | 3172.77777777778
4546.7778 |  | BLUE WELL #2 | 4546.77777777778
3787.7778 |  | BLUE WELL #3 | 3787.77777777778
STANDARD #9
PHOTO #3
3804 |  | RED WELL #1 | 3804
5936 |  | RED WELL #2 | 5936
5450 |  | RED WELL #3 | 5450
9747.5 |  | GREEN WELL #1 | 9747.5
12153.5 |  | GREEN WELL #2 | 12153.5
10724.5 |  | GREEN WELL #3 | 10724.5
6538.7778 |  | BLUE WELL #1 | 6538.77777777778
8043.7778 |  | BLUE WELL #2 | 8043.77777777778
7205.7778 |  | BLUE WELL #3 | 7205.77777777778
|  | POST CALIBRATION | STANDARD #1
|  | PHOTO #1 | GREEN 16X
-2.5097 |  | RED WELL #1 | -2.50966309070236
-0.38541 |  | RED WELL #2 | -0.385408022098955
-7.5724 |  | RED WELL #3 | -7.57235817032933
31576.5915 |  | GREEN WELL #1 | 31576.5915259897
32658.5511 |  | GREEN WELL #2 | 32658.5510623826
30237.568 |  | GREEN WELL #3 | 30237.5680036929
242927.5556 |  | BLUE WELL #1 | 242927.555555556
277211.6667 |  | BLUE WELL #2 | 277211.666666667
238783.2222 |  | BLUE WELL #3 | 238783.222222222
STANDARD #1
PHOTO #2
7.4041 |  | RED WELL #1 | 7.40409838728264
6.0846 |  | RED WELL #2 | 6.08455446076187
4.914 |  | RED WELL #3 | 4.91403039578079
31616.2598 |  | GREEN WELL #1 | 31616.259797471
32681.6161 |  | GREEN WELL #2 | 32681.6160767248
30288.2878 |  | GREEN WELL #3 | 30288.2877697672
243441.5556 |  | BLUE WELL #1 | 243441.555555556
277549.6667 |  | BLUE WELL #2 | 277549.666666667
238820.2222 |  | BLUE WELL #3 | 238820.222222222
STANDARD #1
PHOTO #3
-4.8944 |  | RED WELL #1 | -4.89443529659404
-5.6991 |  | RED WELL #2 | -5.69914643860366
2.6583 |  | RED WELL #3 | 2.65832777455564
31630.1049 |  | GREEN WELL #1 | 31630.1048574495
32681.474 |  | GREEN WELL #2 | 32681.4739882708
30282.656 |  | GREEN WELL #3 | 30282.6559862646
243415.0556 |  | BLUE WELL #1 | 243415.055555556
277813.1667 |  | BLUE WELL #2 | 277813.166666667
238822.7222 |  | BLUE WELL #3 | 238822.722222222
|  | POST CALIBRATION | STANDARD #2
|  | PHOTO #1 | GREEN 32X
17.8785 |  | RED WELL #1 | 17.8785015033291
-46.9529 |  | RED WELL #2 | -46.952919564003
-6.6302 |  | RED WELL #3 | -6.63018297018718
16136.7796 |  | GREEN WELL #1 | 16136.7795779846
16504.2843 |  | GREEN WELL #2 | 16504.2843352677
15041.7743 |  | GREEN WELL #3 | 15041.7742539577
122772.0556 |  | BLUE WELL #1 | 122772.055555556
138384.1667 |  | BLUE WELL #2 | 138384.166666667
117637.7222 |  | BLUE WELL #3 | 117637.722222222
STANDARD #2
PHOTO #2
16.0044 |  | RED WELL #1 | 16.0044390070101
-37.7215 |  | RED WELL #2 | -37.7214636950906
5.4115 |  | RED WELL #3 | 5.41145153159309
16150.9608 |  | GREEN WELL #1 | 16150.9608249337
16511.0629 |  | GREEN WELL #2 | 16511.0628603157
15051.0792 |  | GREEN WELL #3 | 15051.0791958411
122855.0556 |  | BLUE WELL #1 | 122855.055555556
138432.1667 |  | BLUE WELL #2 | 138432.166666667
117481.7222 |  | BLUE WELL #3 | 117481.722222222
STANDARD #2
PHOTO #3
16.4855 |  | RED WELL #1 | 16.4854969863315
-39.7586 |  | RED WELL #2 | -39.7585833778364
-19.1028 |  | RED WELL #3 | -19.1028318349641
16164.9817 |  | GREEN WELL #1 | 16164.9817135739
16514.7433 |  | GREEN WELL #2 | 16514.7433467568
15061.8909 |  | GREEN WELL #3 | 15061.8909094246
123064.0556 |  | BLUE WELL #1 | 123064.055555556
138538.1667 |  | BLUE WELL #2 | 138538.166666667
117617.7222 |  | BLUE WELL #3 | 117617.722222222
|  | POST CALIBRATION | STANDARD #3
|  | PHOTO #1 | GREEN 64X
36.5657 |  | RED WELL #1 | 36.5657156266077
3.3174 |  | RED WELL #2 | 3.31742873649622
18.3284 |  | RED WELL #3 | 18.3284026211506
6944.4105 |  | GREEN WELL #1 | 6944.41051548252
7689.1675 |  | GREEN WELL #2 | 7689.16752765322
6965.1258 |  | GREEN WELL #3 | 6965.12582163531
53923.0556 |  | BLUE WELL #1 | 53923.0555555555
64817.1667 |  | BLUE WELL #2 | 64817.1666666667
55228.7222 |  | BLUE WELL #3 | 55228.7222222222
STANDARD #3
PHOTO #2
33.8929 |  | RED WELL #1 | 33.8928506632424
6.9232 |  | RED WELL #2 | 6.92323041973007
20.4556 |  | RED WELL #3 | 20.4555590033978
6963.9037 |  | GREEN WELL #1 | 6963.90368720316
7693.7058 |  | GREEN WELL #2 | 7693.70580738733
6985.6622 |  | GREEN WELL #3 | 6985.66223693872
53949.0556 |  | BLUE WELL #1 | 53949.0555555555
64905.1667 |  | BLUE WELL #2 | 64905.1666666667
55254.7222 |  | BLUE WELL #3 | 55254.7222222222
STANDARD #3
PHOTO #3
43.5449 |  | RED WELL #1 | 43.5448731783933
8.4658 |  | RED WELL #2 | 8.46583303967466
21.7741 |  | RED WELL #3 | 21.7741362049694
6959.3053 |  | GREEN WELL #1 | 6959.30531089532
7692.4054 |  | GREEN WELL #2 | 7692.40541539693
6979.7567 |  | GREEN WELL #3 | 6979.75667352731
54030.0556 |  | BLUE WELL #1 | 54030.0555555555
64806.1667 |  | BLUE WELL #2 | 64806.1666666667
55307.7222 |  | BLUE WELL #3 | 55307.7222222222
|  | POST CALIBRATION | STANDARD #4
|  | PHOTO #1 | RED 16X
30396.6914 |  | RED WELL #1 | 30396.6914389833
33805.284 |  | RED WELL #2 | 33805.2839575312
30135.6806 |  | RED WELL #3 | 30135.6806050689
6.2552 |  | GREEN WELL #1 | 6.255228312571
-1.6774 |  | GREEN WELL #2 | -1.67736862025563
-0.28067 |  | GREEN WELL #3 | -0.280673812642026
39230.0556 |  | BLUE WELL #1 | 39230.0555555555
39680.1667 |  | BLUE WELL #2 | 39680.1666666667
35066.7222 |  | BLUE WELL #3 | 35066.7222222222
STANDARD #4
PHOTO #2
30395.9644 |  | RED WELL #1 | 30395.9644040091
33816.328 |  | RED WELL #2 | 33816.3280068033
30126.7992 |  | RED WELL #3 | 30126.7992301837
-5.578 |  | GREEN WELL #1 | -5.57800788110962
-2.4662 |  | GREEN WELL #2 | -2.46620934247507
0.48923 |  | GREEN WELL #3 | 0.489230315363524
39136.5556 |  | BLUE WELL #1 | 39136.5555555555
39699.6667 |  | BLUE WELL #2 | 39699.6666666667
35043.2222 |  | BLUE WELL #3 | 35043.2222222222
STANDARD #4
PHOTO #3
30392.4504 |  | RED WELL #1 | 30392.4504209291
33783.9863 |  | RED WELL #2 | 33783.9863175273
30131.3865 |  | RED WELL #3 | 30131.3864851251
-0.67722 |  | GREEN WELL #1 | -0.67722043146016
4.1436 |  | GREEN WELL #2 | 4.14357796273687
-0.20856 |  | GREEN WELL #3 | -0.208556502717991
38936.0556 |  | BLUE WELL #1 | 38936.0555555555
39773.1667 |  | BLUE WELL #2 | 39773.1666666667
34956.7222 |  | BLUE WELL #3 | 34956.7222222222
|  | POST CALIBRATION | STANDARD #5
|  | PHOTO #1 | RED 32X
14129.5469 |  | RED WELL #1 | 14129.5468779307
16015.1264 |  | RED WELL #2 | 16015.1264111177
13988.94 |  | RED WELL #3 | 13988.9400205013
-12.9887 |  | GREEN WELL #1 | -12.9886791902601
-16.4329 |  | GREEN WELL #2 | -16.4329360732366
-37.012 |  | GREEN WELL #3 | -37.0119775658153
18430.5556 |  | BLUE WELL #1 | 18430.5555555556
18993.6667 |  | BLUE WELL #2 | 18993.6666666667
16411.2222 |  | BLUE WELL #3 | 16411.2222222222
STANDARD #5
PHOTO #2
14141.8121 |  | RED WELL #1 | 14141.8121019555
16026.0289 |  | RED WELL #2 | 16026.0289349311
14001.4313 |  | RED WELL #3 | 14001.4312686767
-10.6326 |  | GREEN WELL #1 | -10.6325797121323
-15.038 |  | GREEN WELL #2 | -15.0380241465033
-29.9419 |  | GREEN WELL #3 | -29.9418906605273
18412.5556 |  | BLUE WELL #1 | 18412.5555555556
18846.6667 |  | BLUE WELL #2 | 18846.6666666667
16401.2222 |  | BLUE WELL #3 | 16401.2222222222
STANDARD #5
PHOTO #3
14126.8329 |  | RED WELL #1 | 14126.8328774416
16011.8442 |  | RED WELL #2 | 16011.8442334893
13976.46 |  | RED WELL #3 | 13976.4600297558
-6.678 |  | GREEN WELL #1 | -6.67798845899011
-9.6346 |  | GREEN WELL #2 | -9.63455649212314
-30.1524 |  | GREEN WELL #3 | -30.1523745521168
18387.5556 |  | BLUE WELL #1 | 18387.5555555556
18708.6667 |  | BLUE WELL #2 | 18708.6666666667
16307.2222 |  | BLUE WELL #3 | 16307.2222222222
|  | POST CALIBRATION | STANDARD #6
|  | PHOTO #1 | RED 64X
6656.216 |  | RED WELL #1 | 6656.21598031865
6491.2057 |  | RED WELL #2 | 6491.20570324271
6280.4589 |  | RED WELL #3 | 6280.45885479089
-37.0588 |  | GREEN WELL #1 | -37.0588288685413
-11.2977 |  | GREEN WELL #2 | -11.2977141359592
-24.1074 |  | GREEN WELL #3 | -24.107355672117
9429.5556 |  | BLUE WELL #1 | 9429.55555555555
8606.6667 |  | BLUE WELL #2 | 8606.66666666667
8201.2222 |  | BLUE WELL #3 | 8201.22222222222
STANDARD #6
PHOTO #2
6646.3781 |  | RED WELL #1 | 6646.37805868091
6496.0189 |  | RED WELL #2 | 6496.01890389183
6283.1127 |  | RED WELL #3 | 6283.11271857649
-30.7288 |  | GREEN WELL #1 | -30.7288335248728
-8.9411 |  | GREEN WELL #2 | -8.94114902408104
-16.054 |  | GREEN WELL #3 | -16.053964491356
9578.5556 |  | BLUE WELL #1 | 9578.55555555555
8566.6667 |  | BLUE WELL #2 | 8566.66666666667
8146.2222 |  | BLUE WELL #3 | 8146.22222222222
STANDARD #6
PHOTO #3
6646.9031 |  | RED WELL #1 | 6646.9031018602
6501.9354 |  | RED WELL #2 | 6501.93542054259
6269.4009 |  | RED WELL #3 | 6269.40085646978
-26.6793 |  | GREEN WELL #1 | -26.6793365620442
-6.4966 |  | GREEN WELL #2 | -6.49659519787038
-12.8854 |  | GREEN WELL #3 | -12.8853904211672
9470.0556 |  | BLUE WELL #1 | 9470.05555555555
8632.1667 |  | BLUE WELL #2 | 8632.16666666667
8355.7222 |  | BLUE WELL #3 | 8355.72222222222
|  | POST CALIBRATION | STANDARD #7
|  | PHOTO #1 | BUFFER #1
4.2951 |  | RED WELL #1 | 4.29507201011951
-10.1115 |  | RED WELL #2 | -10.1114627286696
7.9328 |  | RED WELL #3 | 7.93276652418206
-16.7071 |  | GREEN WELL #1 | -16.7071142246698
-8.2873 |  | GREEN WELL #2 | -8.28730160416204
-8.8929 |  | GREEN WELL #3 | -8.89285131595874
160.0556 |  | BLUE WELL #1 | 160.055555555556
160.1667 |  | BLUE WELL #2 | 160.166666666667
253.7222 |  | BLUE WELL #3 | 253.722222222223
STANDARD #7
PHOTO #2
8.8216 |  | RED WELL #1 | 8.82156687743035
-3.4829 |  | RED WELL #2 | -3.48288922911448
12.1115 |  | RED WELL #3 | 12.111466600377
-14.8484 |  | GREEN WELL #1 | -14.8483627479272
-7.2761 |  | GREEN WELL #2 | -7.27607232956695
-4.3011 |  | GREEN WELL #3 | -4.30105102022712
35.0556 |  | BLUE WELL #1 | 35.0555555555557
202.1667 |  | BLUE WELL #2 | 202.166666666667
290.7222 |  | BLUE WELL #3 | 290.722222222223
STANDARD #7
PHOTO #3
10.0185 |  | RED WELL #1 | 10.0184802192698
-6.4329 |  | RED WELL #2 | -6.43285913620193
9.761 |  | RED WELL #3 | 9.76099582012333
-17.2251 |  | GREEN WELL #1 | -17.2250973021501
-4.7963 |  | GREEN WELL #2 | -4.79627084421514
-6.7195 |  | GREEN WELL #3 | -6.71949522973905
187.0556 |  | BLUE WELL #1 | 187.055555555556
324.1667 |  | BLUE WELL #2 | 324.166666666667
297.7222 |  | BLUE WELL #3 | 297.722222222223
|  | POST CALIBRATION | STANDARD #8
|  | PHOTO #1 | BUFFER #2
44.6902 |  | RED WELL #1 | 44.6902417262101
21.0249 |  | RED WELL #2 | 21.0248757976513
3.0759 |  | RED WELL #3 | 3.07591764414768
22.9856 |  | GREEN WELL #1 | 22.9855583372967
8.0135 |  | GREEN WELL #2 | 8.0134856653508
5.9713 |  | GREEN WELL #3 | 5.97132450590768
680.0556 |  | BLUE WELL #1 | 680.055555555556
431.1667 |  | BLUE WELL #2 | 431.166666666667
410.7222 |  | BLUE WELL #3 | 410.722222222223
STANDARD #8
PHOTO #2
-32.0434 |  | RED WELL #1 | -32.0433563270785
-61.7596 |  | RED WELL #2 | -61.7596227360597
-79.998 |  | RED WELL #3 | -79.9979541588013
5.9546 |  | GREEN WELL #1 | 5.95457516386857
-9.1977 |  | GREEN WELL #2 | -9.19769823845091
-5.9472 |  | GREEN WELL #3 | -5.94718703455439
-686.4444 |  | BLUE WELL #1 | -686.444444444444
-974.3333 |  | BLUE WELL #2 | -974.333333333333
-972.7778 |  | BLUE WELL #3 | -972.777777777777
STANDARD #8
PHOTO #3
-23.1424 |  | RED WELL #1 | -23.1423623328577
-55.5204 |  | RED WELL #2 | -55.5204417953464
-75.2525 |  | RED WELL #3 | -75.2525006813337
13.2131 |  | GREEN WELL #1 | 13.213106818688
-7.4665 |  | GREEN WELL #2 | -7.46650267502834
-7.4589 |  | GREEN WELL #3 | -7.45888527438266
-646.9444 |  | BLUE WELL #1 | -646.944444444444
-802.8333 |  | BLUE WELL #2 | -802.833333333333
-848.2778 |  | BLUE WELL #3 | -848.277777777777
|  | POST CALIBRATION | STANDARD #9
|  | PHOTO #1 | BUFFER #3
-44.0365 |  | RED WELL #1 | -44.0364906144248
-4.109 |  | RED WELL #2 | -4.10902507477282
1.1314 |  | RED WELL #3 | 1.13135119744173
24.1128 |  | GREEN WELL #1 | 24.112754671342
30.6577 |  | GREEN WELL #2 | 30.6577157325565
30.7965 |  | GREEN WELL #3 | 30.7964688782677
588.0556 |  | BLUE WELL #1 | 588.055555555556
693.1667 |  | BLUE WELL #2 | 693.166666666667
603.7222 |  | BLUE WELL #3 | 603.722222222223
STANDARD #9
PHOTO #2
57.3167 |  | RED WELL #1 | 57.3167398551136
101.7249 |  | RED WELL #2 | 101.724905478558
96.3312 |  | RED WELL #3 | 96.3311899869857
-13.3092 |  | GREEN WELL #1 | -13.3092450307394
-4.5507 |  | GREEN WELL #2 | -4.55073563982736
-6.4675 |  | GREEN WELL #3 | -6.46747916444341
-836.4444 |  | BLUE WELL #1 | -836.444444444444
-760.3333 |  | BLUE WELL #2 | -760.333333333333
-721.7778 |  | BLUE WELL #3 | -721.777777777777
STANDARD #9
PHOTO #3
-25.9199 |  | RED WELL #1 | -25.9198914137822
18.6665 |  | RED WELL #2 | 18.666519423956
24.9068 |  | RED WELL #3 | 24.9067670668778
-4.1762 |  | GREEN WELL #1 | -4.17617568570892
2.9034 |  | GREEN WELL #2 | 2.90337993334345
3.0192 |  | GREEN WELL #3 | 3.01915565512972
519.5556 |  | BLUE WELL #1 | 519.555555555556
726.6667 |  | BLUE WELL #2 | 726.666666666667
686.2222 |  | BLUE WELL #3 | 686.222222222223

## REF Cube Readings
REF Cube Host Name (auto populated) | B19460006
STANDARD #1
GREEN 16X
-111.2651 |  | RED WELL #1 | -111.265132095765
-156.1821 |  | RED WELL #2 | -156.182144192096
-148.6273 |  | RED WELL #3 | -148.627255243961
31552.1055 |  | GREEN WELL #1 | 31552.1054876458
32675.1877 |  | GREEN WELL #2 | 32675.1876734995
30245.7547 |  | GREEN WELL #3 | 30245.7546955734
STANDARD #2
GREEN 32X
-85.5476 |  | RED WELL #1 | -85.547595287457
-212.5217 |  | RED WELL #2 | -212.521693575218
-257.7427 |  | RED WELL #3 | -257.742674521806
16213.0861 |  | GREEN WELL #1 | 16213.0861281166
16322.6945 |  | GREEN WELL #2 | 16322.6944738129
15356.9713 |  | GREEN WELL #3 | 15356.971260254
STANDARD #3
GREEN 64X
-111.2139 |  | RED WELL #1 | -111.213891544751
-229.4637 |  | RED WELL #2 | -229.463692519956
-300.6165 |  | RED WELL #3 | -300.616457566443
6727.4892 |  | GREEN WELL #1 | 6727.48919600178
7739.0747 |  | GREEN WELL #2 | 7739.07465793084
7202.2284 |  | GREEN WELL #3 | 7202.22838999478
STANDARD #4
RED 16X
30432.9104 |  | RED WELL #1 | 30432.9104142137
33696.0392 |  | RED WELL #2 | 33696.0392275006
30077.4416 |  | RED WELL #3 | 30077.4416083016
-237.5394 |  | GREEN WELL #1 | -237.539370565846
-211.744 |  | GREEN WELL #2 | -211.743998133493
-290.8598 |  | GREEN WELL #3 | -290.859833813174
STANDARD #5
RED 32X
14191.0297 |  | RED WELL #1 | 14191.0297436141
16213.1043 |  | RED WELL #2 | 16213.1043268878
13806.2851 |  | RED WELL #3 | 13806.285091622
-175.8102 |  | GREEN WELL #1 | -175.810175061922
-158.6836 |  | GREEN WELL #2 | -158.683629968562
-198.7036 |  | GREEN WELL #3 | -198.703619248004
STANDARD #6
RED 64X
6598.9444 |  | RED WELL #1 | 6598.94442213181
6094.7451 |  | RED WELL #2 | 6094.74510256631
6107.8198 |  | RED WELL #3 | 6107.81976463659
-169.1524 |  | GREEN WELL #1 | -169.152405046869
-213.638 |  | GREEN WELL #2 | -213.637969754123
-91.553 |  | GREEN WELL #3 | -91.5530016409894
STANDARD #7
BUFFER #1
-135.2569 |  | RED WELL #1 | -135.256859895576
-235.8671 |  | RED WELL #2 | -235.867127440013
-312.4438 |  | RED WELL #3 | -312.443825352633
-117.8029 |  | GREEN WELL #1 | -117.802941534096
-188.6153 |  | GREEN WELL #2 | -188.615313382778
9.4243 |  | GREEN WELL #3 | 9.42430123058784
STANDARD #8
BUFFER #2
-116.4056 |  | RED WELL #1 | -116.40564476188
-214.8377 |  | RED WELL #2 | -214.837657103061
-332.6426 |  | RED WELL #3 | -332.642556021415
-110.7137 |  | GREEN WELL #1 | -110.713714627917
-183.4026 |  | GREEN WELL #2 | -183.402561989712
-11.4689 |  | GREEN WELL #3 | -11.468855439902
STANDARD #9
BUFFER #3
-105.8805 |  | RED WELL #1 | -105.88049951432
-148.319 |  | RED WELL #2 | -148.318985244218
-245.7728 |  | RED WELL #3 | -245.772787058945
-122.1307 |  | GREEN WELL #1 | -122.130728756882
-174.0747 |  | GREEN WELL #2 | -174.074715804937
-4.7726 |  | GREEN WELL #3 | -4.77255950660102

## optical_store
files_1 | standard-1-Reference_1.png | B2030000A
files_2 | standard-1-Reference_2.png
files_3 | standard-1-Reference_3.png
files_4 | standard-2-Reference_1.png
files_5 | standard-2-Reference_2.png
files_6 | standard-2-Reference_3.png
files_7 | standard-3-Reference_1.png
files_8 | standard-3-Reference_2.png
files_9 | standard-3-Reference_3.png
files_10 | standard-4-Reference_1.png
files_11 | standard-4-Reference_2.png
files_12 | standard-4-Reference_3.png
files_13 | standard-5-Reference_1.png
files_14 | standard-5-Reference_2.png
files_15 | standard-5-Reference_3.png
files_16 | standard-6-Reference_1.png
files_17 | standard-6-Reference_2.png
files_18 | standard-6-Reference_3.png
files_19 | standard-7-Reference_1.png
files_20 | standard-7-Reference_2.png
files_21 | standard-7-Reference_3.png
files_22 | standard-8-Reference_1.png
files_23 | standard-8-Reference_2.png
files_24 | standard-8-Reference_3.png
files_25 | standard-9-Reference_1.png
files_26 | standard-9-Reference_2.png
files_27 | standard-9-Reference_3.png
files_28 | /nfs/r3_mes_data/optical_cal_targets_gen_3_0.json
test_results_analysis_values_precal_standard_1_img1_1 | 146962
test_results_analysis_values_precal_standard_1_img1_2 | 169754
test_results_analysis_values_precal_standard_1_img1_3 | 141195
test_results_analysis_values_precal_standard_1_img1_4 | 766650.5
test_results_analysis_values_precal_standard_1_img1_5 | 939531.5
test_results_analysis_values_precal_standard_1_img1_6 | 768666.5
test_results_analysis_values_precal_standard_1_img1_7 | 254859.7778
test_results_analysis_values_precal_standard_1_img1_8 | 290441.7778
test_results_analysis_values_precal_standard_1_img1_9 | 251215.7778
test_results_analysis_values_precal_standard_1_img2_1 | 147634
test_results_analysis_values_precal_standard_1_img2_2 | 170300
test_results_analysis_values_precal_standard_1_img2_3 | 141951
test_results_analysis_values_precal_standard_1_img2_4 | 767902.5
test_results_analysis_values_precal_standard_1_img2_5 | 940484.5
test_results_analysis_values_precal_standard_1_img2_6 | 770237.5
test_results_analysis_values_precal_standard_1_img2_7 | 255529.7778
test_results_analysis_values_precal_standard_1_img2_8 | 290935.7778
test_results_analysis_values_precal_standard_1_img2_9 | 251408.7778
test_results_analysis_values_precal_standard_1_img3_1 | 147542
test_results_analysis_values_precal_standard_1_img3_2 | 170169
test_results_analysis_values_precal_standard_1_img3_3 | 141942
test_results_analysis_values_precal_standard_1_img3_4 | 768204.5
test_results_analysis_values_precal_standard_1_img3_5 | 940463.5
test_results_analysis_values_precal_standard_1_img3_6 | 770108.5
test_results_analysis_values_precal_standard_1_img3_7 | 255767.7778
test_results_analysis_values_precal_standard_1_img3_8 | 291463.7778
test_results_analysis_values_precal_standard_1_img3_9 | 251675.7778
test_results_analysis_values_precal_standard_2_img1_1 | 77758
test_results_analysis_values_precal_standard_2_img1_2 | 88130
test_results_analysis_values_precal_standard_2_img1_3 | 73036
test_results_analysis_values_precal_standard_2_img1_4 | 397115.5
test_results_analysis_values_precal_standard_2_img1_5 | 481130.5
test_results_analysis_values_precal_standard_2_img1_6 | 388330.5
test_results_analysis_values_precal_standard_2_img1_7 | 131288.7778
test_results_analysis_values_precal_standard_2_img1_8 | 148198.7778
test_results_analysis_values_precal_standard_2_img1_9 | 126654.7778
test_results_analysis_values_precal_standard_2_img2_1 | 77743
test_results_analysis_values_precal_standard_2_img2_2 | 88260
test_results_analysis_values_precal_standard_2_img2_3 | 73225
test_results_analysis_values_precal_standard_2_img2_4 | 397608.5
test_results_analysis_values_precal_standard_2_img2_5 | 481511.5
test_results_analysis_values_precal_standard_2_img2_6 | 388759.5
test_results_analysis_values_precal_standard_2_img2_7 | 131526.7778
test_results_analysis_values_precal_standard_2_img2_8 | 148401.7778
test_results_analysis_values_precal_standard_2_img2_9 | 126653.7778
test_results_analysis_values_precal_standard_2_img3_1 | 78073
test_results_analysis_values_precal_standard_2_img3_2 | 88507
test_results_analysis_values_precal_standard_2_img3_3 | 73137
test_results_analysis_values_precal_standard_2_img3_4 | 398044.5
test_results_analysis_values_precal_standard_2_img3_5 | 481714.5
test_results_analysis_values_precal_standard_2_img3_6 | 389060.5
test_results_analysis_values_precal_standard_2_img3_7 | 131864.7778
test_results_analysis_values_precal_standard_2_img3_8 | 148636.7778
test_results_analysis_values_precal_standard_2_img3_9 | 126918.7778
test_results_analysis_values_precal_standard_3_img1_1 | 36477
test_results_analysis_values_precal_standard_3_img1_2 | 44481
test_results_analysis_values_precal_standard_3_img1_3 | 36899
test_results_analysis_values_precal_standard_3_img1_4 | 176285.5
test_results_analysis_values_precal_standard_3_img1_5 | 229967.5
test_results_analysis_values_precal_standard_3_img1_6 | 184908.5
test_results_analysis_values_precal_standard_3_img1_7 | 59898.7778
test_results_analysis_values_precal_standard_3_img1_8 | 72090.7778
test_results_analysis_values_precal_standard_3_img1_9 | 61704.7778
test_results_analysis_values_precal_standard_3_img2_1 | 36690
test_results_analysis_values_precal_standard_3_img2_2 | 44731
test_results_analysis_values_precal_standard_3_img2_3 | 37195
test_results_analysis_values_precal_standard_3_img2_4 | 176840.5
test_results_analysis_values_precal_standard_3_img2_5 | 230210.5
test_results_analysis_values_precal_standard_3_img2_6 | 185524.5
test_results_analysis_values_precal_standard_3_img2_7 | 60226.7778
test_results_analysis_values_precal_standard_3_img2_8 | 72480.7778
test_results_analysis_values_precal_standard_3_img2_9 | 62032.7778
test_results_analysis_values_precal_standard_3_img3_1 | 36933
test_results_analysis_values_precal_standard_3_img3_2 | 44851
test_results_analysis_values_precal_standard_3_img3_3 | 37293
test_results_analysis_values_precal_standard_3_img3_4 | 176868.5
test_results_analysis_values_precal_standard_3_img3_5 | 230284.5
test_results_analysis_values_precal_standard_3_img3_6 | 185489.5
test_results_analysis_values_precal_standard_3_img3_7 | 60420.7778
test_results_analysis_values_precal_standard_3_img3_8 | 72494.7778
test_results_analysis_values_precal_standard_3_img3_9 | 62198.7778
test_results_analysis_values_precal_standard_4_img1_1 | 527259
test_results_analysis_values_precal_standard_4_img1_2 | 545618
test_results_analysis_values_precal_standard_4_img1_3 | 508473
test_results_analysis_values_precal_standard_4_img1_4 | 106254.5
test_results_analysis_values_precal_standard_4_img1_5 | 107835.5
test_results_analysis_values_precal_standard_4_img1_6 | 99182.5
test_results_analysis_values_precal_standard_4_img1_7 | 45034.7778
test_results_analysis_values_precal_standard_4_img1_8 | 46782.7778
test_results_analysis_values_precal_standard_4_img1_9 | 41371.7778
test_results_analysis_values_precal_standard_4_img2_1 | 527531
test_results_analysis_values_precal_standard_4_img2_2 | 546121
test_results_analysis_values_precal_standard_4_img2_3 | 508669
test_results_analysis_values_precal_standard_4_img2_4 | 106331.5
test_results_analysis_values_precal_standard_4_img2_5 | 108200.5
test_results_analysis_values_precal_standard_4_img2_6 | 99532.5
test_results_analysis_values_precal_standard_4_img2_7 | 45206.7778
test_results_analysis_values_precal_standard_4_img2_8 | 47067.7778
test_results_analysis_values_precal_standard_4_img2_9 | 41613.7778
test_results_analysis_values_precal_standard_4_img3_1 | 527501
test_results_analysis_values_precal_standard_4_img3_2 | 545660
test_results_analysis_values_precal_standard_4_img3_3 | 508747
test_results_analysis_values_precal_standard_4_img3_4 | 106472.5
test_results_analysis_values_precal_standard_4_img3_5 | 108332.5
test_results_analysis_values_precal_standard_4_img3_6 | 99565.5
test_results_analysis_values_precal_standard_4_img3_7 | 45111.7778
test_results_analysis_values_precal_standard_4_img3_8 | 47246.7778
test_results_analysis_values_precal_standard_4_img3_9 | 41632.7778
test_results_analysis_values_precal_standard_5_img1_1 | 248184
test_results_analysis_values_precal_standard_5_img1_2 | 262108
test_results_analysis_values_precal_standard_5_img1_3 | 239467
test_results_analysis_values_precal_standard_5_img1_4 | 54060.5
test_results_analysis_values_precal_standard_5_img1_5 | 56684.5
test_results_analysis_values_precal_standard_5_img1_6 | 50536.5
test_results_analysis_values_precal_standard_5_img1_7 | 23795.7778
test_results_analysis_values_precal_standard_5_img1_8 | 25656.7778
test_results_analysis_values_precal_standard_5_img1_9 | 22276.7778
test_results_analysis_values_precal_standard_5_img2_1 | 248569
test_results_analysis_values_precal_standard_5_img2_2 | 262454
test_results_analysis_values_precal_standard_5_img2_3 | 239870
test_results_analysis_values_precal_standard_5_img2_4 | 54189.5
test_results_analysis_values_precal_standard_5_img2_5 | 56789.5
test_results_analysis_values_precal_standard_5_img2_6 | 50781.5
test_results_analysis_values_precal_standard_5_img2_7 | 23942.7778
test_results_analysis_values_precal_standard_5_img2_8 | 25674.7778
test_results_analysis_values_precal_standard_5_img2_9 | 22431.7778
test_results_analysis_values_precal_standard_5_img3_1 | 245855
test_results_analysis_values_precal_standard_5_img3_2 | 259779
test_results_analysis_values_precal_standard_5_img3_3 | 236984
test_results_analysis_values_precal_standard_5_img3_4 | 54353.5
test_results_analysis_values_precal_standard_5_img3_5 | 57018.5
test_results_analysis_values_precal_standard_5_img3_6 | 50822.5
test_results_analysis_values_precal_standard_5_img3_7 | 23924.7778
test_results_analysis_values_precal_standard_5_img3_8 | 25543.7778
test_results_analysis_values_precal_standard_5_img3_9 | 22344.7778
test_results_analysis_values_precal_standard_6_img1_1 | 119276
test_results_analysis_values_precal_standard_6_img1_2 | 110388
test_results_analysis_values_precal_standard_6_img1_3 | 110693
test_results_analysis_values_precal_standard_6_img1_4 | 29980.5
test_results_analysis_values_precal_standard_6_img1_5 | 30042.5
test_results_analysis_values_precal_standard_6_img1_6 | 28343.5
test_results_analysis_values_precal_standard_6_img1_7 | 14590.7778
test_results_analysis_values_precal_standard_6_img1_8 | 15065.7778
test_results_analysis_values_precal_standard_6_img1_9 | 13862.7778
test_results_analysis_values_precal_standard_6_img2_1 | 119416
test_results_analysis_values_precal_standard_6_img2_2 | 110750
test_results_analysis_values_precal_standard_6_img2_3 | 111047
test_results_analysis_values_precal_standard_6_img2_4 | 30201.5
test_results_analysis_values_precal_standard_6_img2_5 | 30224.5
test_results_analysis_values_precal_standard_6_img2_6 | 28651.5
test_results_analysis_values_precal_standard_6_img2_7 | 14897.7778
test_results_analysis_values_precal_standard_6_img2_8 | 15183.7778
test_results_analysis_values_precal_standard_6_img2_9 | 13965.7778
test_results_analysis_values_precal_standard_6_img3_1 | 119582
test_results_analysis_values_precal_standard_6_img3_2 | 110993
test_results_analysis_values_precal_standard_6_img3_3 | 110979
test_results_analysis_values_precal_standard_6_img3_4 | 30319.5
test_results_analysis_values_precal_standard_6_img3_5 | 30330.5
test_results_analysis_values_precal_standard_6_img3_6 | 28711.5
test_results_analysis_values_precal_standard_6_img3_7 | 14908.7778
test_results_analysis_values_precal_standard_6_img3_8 | 15368.7778
test_results_analysis_values_precal_standard_6_img3_9 | 14294.7778
test_results_analysis_values_precal_standard_7_img1_1 | 5504
test_results_analysis_values_precal_standard_7_img1_2 | 6688
test_results_analysis_values_precal_standard_7_img1_3 | 6376
test_results_analysis_values_precal_standard_7_img1_4 | 9862.5
test_results_analysis_values_precal_standard_7_img1_5 | 12077.5
test_results_analysis_values_precal_standard_7_img1_6 | 10699.5
test_results_analysis_values_precal_standard_7_img1_7 | 5480.7778
test_results_analysis_values_precal_standard_7_img1_8 | 6778.7778
test_results_analysis_values_precal_standard_7_img1_9 | 6074.7778
test_results_analysis_values_precal_standard_7_img2_1 | 5847
test_results_analysis_values_precal_standard_7_img2_2 | 7055
test_results_analysis_values_precal_standard_7_img2_3 | 6723
test_results_analysis_values_precal_standard_7_img2_4 | 10006.5
test_results_analysis_values_precal_standard_7_img2_5 | 12210.5
test_results_analysis_values_precal_standard_7_img2_6 | 10910.5
test_results_analysis_values_precal_standard_7_img2_7 | 5721.7778
test_results_analysis_values_precal_standard_7_img2_8 | 7186.7778
test_results_analysis_values_precal_standard_7_img2_9 | 6477.7778
test_results_analysis_values_precal_standard_7_img3_1 | 5932
test_results_analysis_values_precal_standard_7_img3_2 | 7097
test_results_analysis_values_precal_standard_7_img3_3 | 6750
test_results_analysis_values_precal_standard_7_img3_4 | 9977.5
test_results_analysis_values_precal_standard_7_img3_5 | 12294.5
test_results_analysis_values_precal_standard_7_img3_6 | 10867.5
test_results_analysis_values_precal_standard_7_img3_7 | 5849.7778
test_results_analysis_values_precal_standard_7_img3_8 | 7284.7778
test_results_analysis_values_precal_standard_7_img3_9 | 6460.7778
test_results_analysis_values_precal_standard_8_img1_1 | 6492
test_results_analysis_values_precal_standard_8_img1_2 | 7391
test_results_analysis_values_precal_standard_8_img1_3 | 6504
test_results_analysis_values_precal_standard_8_img1_4 | 11014.5
test_results_analysis_values_precal_standard_8_img1_5 | 12715.5
test_results_analysis_values_precal_standard_8_img1_6 | 11147.5
test_results_analysis_values_precal_standard_8_img1_7 | 6254.7778
test_results_analysis_values_precal_standard_8_img1_8 | 7303.7778
test_results_analysis_values_precal_standard_8_img1_9 | 6485.7778
test_results_analysis_values_precal_standard_8_img2_1 | 3660
test_results_analysis_values_precal_standard_8_img2_2 | 4546
test_results_analysis_values_precal_standard_8_img2_3 | 3631
test_results_analysis_values_precal_standard_8_img2_4 | 9874.5
test_results_analysis_values_precal_standard_8_img2_5 | 11502.5
test_results_analysis_values_precal_standard_8_img2_6 | 10113.5
test_results_analysis_values_precal_standard_8_img2_7 | 3302.7778
test_results_analysis_values_precal_standard_8_img2_8 | 4312.7778
test_results_analysis_values_precal_standard_8_img2_9 | 3516.7778
test_results_analysis_values_precal_standard_8_img3_1 | 3903
test_results_analysis_values_precal_standard_8_img3_2 | 4714
test_results_analysis_values_precal_standard_8_img3_3 | 3764
test_results_analysis_values_precal_standard_8_img3_4 | 10142.5
test_results_analysis_values_precal_standard_8_img3_5 | 11638.5
test_results_analysis_values_precal_standard_8_img3_6 | 10160.5
test_results_analysis_values_precal_standard_8_img3_7 | 3452.7778
test_results_analysis_values_precal_standard_8_img3_8 | 4594.7778
test_results_analysis_values_precal_standard_8_img3_9 | 3751.7778
test_results_analysis_values_precal_standard_9_img1_1 | 3342
test_results_analysis_values_precal_standard_9_img1_2 | 5435
test_results_analysis_values_precal_standard_9_img1_3 | 4903
test_results_analysis_values_precal_standard_9_img1_4 | 10722.5
test_results_analysis_values_precal_standard_9_img1_5 | 13232.5
test_results_analysis_values_precal_standard_9_img1_6 | 11705.5
test_results_analysis_values_precal_standard_9_img1_7 | 6066.7778
test_results_analysis_values_precal_standard_9_img1_8 | 7469.7778
test_results_analysis_values_precal_standard_9_img1_9 | 6582.7778
test_results_analysis_values_precal_standard_9_img2_1 | 6973
test_results_analysis_values_precal_standard_9_img2_2 | 9008
test_results_analysis_values_precal_standard_9_img2_3 | 8385
test_results_analysis_values_precal_standard_9_img2_4 | 9684.5
test_results_analysis_values_precal_standard_9_img2_5 | 12069.5
test_results_analysis_values_precal_standard_9_img2_6 | 10590.5
test_results_analysis_values_precal_standard_9_img2_7 | 3172.7778
test_results_analysis_values_precal_standard_9_img2_8 | 4546.7778
test_results_analysis_values_precal_standard_9_img2_9 | 3787.7778
test_results_analysis_values_precal_standard_9_img3_1 | 3804
test_results_analysis_values_precal_standard_9_img3_2 | 5936
test_results_analysis_values_precal_standard_9_img3_3 | 5450
test_results_analysis_values_precal_standard_9_img3_4 | 9747.5
test_results_analysis_values_precal_standard_9_img3_5 | 12153.5
test_results_analysis_values_precal_standard_9_img3_6 | 10724.5
test_results_analysis_values_precal_standard_9_img3_7 | 6538.7778
test_results_analysis_values_precal_standard_9_img3_8 | 8043.7778
test_results_analysis_values_precal_standard_9_img3_9 | 7205.7778
test_results_analysis_values_postcal_standard_1_img1_1 | -2.5097
test_results_analysis_values_postcal_standard_1_img1_2 | -0.38541
test_results_analysis_values_postcal_standard_1_img1_3 | -7.5724
test_results_analysis_values_postcal_standard_1_img1_4 | 31576.5915
test_results_analysis_values_postcal_standard_1_img1_5 | 32658.5511
test_results_analysis_values_postcal_standard_1_img1_6 | 30237.568
test_results_analysis_values_postcal_standard_1_img1_7 | 242927.5556
test_results_analysis_values_postcal_standard_1_img1_8 | 277211.6667
test_results_analysis_values_postcal_standard_1_img1_9 | 238783.2222
test_results_analysis_values_postcal_standard_1_img2_1 | 7.4041
test_results_analysis_values_postcal_standard_1_img2_2 | 6.0846
test_results_analysis_values_postcal_standard_1_img2_3 | 4.914
test_results_analysis_values_postcal_standard_1_img2_4 | 31616.2598
test_results_analysis_values_postcal_standard_1_img2_5 | 32681.6161
test_results_analysis_values_postcal_standard_1_img2_6 | 30288.2878
test_results_analysis_values_postcal_standard_1_img2_7 | 243441.5556
test_results_analysis_values_postcal_standard_1_img2_8 | 277549.6667
test_results_analysis_values_postcal_standard_1_img2_9 | 238820.2222
test_results_analysis_values_postcal_standard_1_img3_1 | -4.8944
test_results_analysis_values_postcal_standard_1_img3_2 | -5.6991
test_results_analysis_values_postcal_standard_1_img3_3 | 2.6583
test_results_analysis_values_postcal_standard_1_img3_4 | 31630.1049
test_results_analysis_values_postcal_standard_1_img3_5 | 32681.474
test_results_analysis_values_postcal_standard_1_img3_6 | 30282.656
test_results_analysis_values_postcal_standard_1_img3_7 | 243415.0556
test_results_analysis_values_postcal_standard_1_img3_8 | 277813.1667
test_results_analysis_values_postcal_standard_1_img3_9 | 238822.7222
test_results_analysis_values_postcal_standard_2_img1_1 | 17.8785
test_results_analysis_values_postcal_standard_2_img1_2 | -46.9529
test_results_analysis_values_postcal_standard_2_img1_3 | -6.6302
test_results_analysis_values_postcal_standard_2_img1_4 | 16136.7796
test_results_analysis_values_postcal_standard_2_img1_5 | 16504.2843
test_results_analysis_values_postcal_standard_2_img1_6 | 15041.7743
test_results_analysis_values_postcal_standard_2_img1_7 | 122772.0556
test_results_analysis_values_postcal_standard_2_img1_8 | 138384.1667
test_results_analysis_values_postcal_standard_2_img1_9 | 117637.7222
test_results_analysis_values_postcal_standard_2_img2_1 | 16.0044
test_results_analysis_values_postcal_standard_2_img2_2 | -37.7215
test_results_analysis_values_postcal_standard_2_img2_3 | 5.4115
test_results_analysis_values_postcal_standard_2_img2_4 | 16150.9608
test_results_analysis_values_postcal_standard_2_img2_5 | 16511.0629
test_results_analysis_values_postcal_standard_2_img2_6 | 15051.0792
test_results_analysis_values_postcal_standard_2_img2_7 | 122855.0556
test_results_analysis_values_postcal_standard_2_img2_8 | 138432.1667
test_results_analysis_values_postcal_standard_2_img2_9 | 117481.7222
test_results_analysis_values_postcal_standard_2_img3_1 | 16.4855
test_results_analysis_values_postcal_standard_2_img3_2 | -39.7586
test_results_analysis_values_postcal_standard_2_img3_3 | -19.1028
test_results_analysis_values_postcal_standard_2_img3_4 | 16164.9817
test_results_analysis_values_postcal_standard_2_img3_5 | 16514.7433
test_results_analysis_values_postcal_standard_2_img3_6 | 15061.8909
test_results_analysis_values_postcal_standard_2_img3_7 | 123064.0556
test_results_analysis_values_postcal_standard_2_img3_8 | 138538.1667
test_results_analysis_values_postcal_standard_2_img3_9 | 117617.7222
test_results_analysis_values_postcal_standard_3_img1_1 | 36.5657
test_results_analysis_values_postcal_standard_3_img1_2 | 3.3174
test_results_analysis_values_postcal_standard_3_img1_3 | 18.3284
test_results_analysis_values_postcal_standard_3_img1_4 | 6944.4105
test_results_analysis_values_postcal_standard_3_img1_5 | 7689.1675
test_results_analysis_values_postcal_standard_3_img1_6 | 6965.1258
test_results_analysis_values_postcal_standard_3_img1_7 | 53923.0556
test_results_analysis_values_postcal_standard_3_img1_8 | 64817.1667
test_results_analysis_values_postcal_standard_3_img1_9 | 55228.7222
test_results_analysis_values_postcal_standard_3_img2_1 | 33.8929
test_results_analysis_values_postcal_standard_3_img2_2 | 6.9232
test_results_analysis_values_postcal_standard_3_img2_3 | 20.4556
test_results_analysis_values_postcal_standard_3_img2_4 | 6963.9037
test_results_analysis_values_postcal_standard_3_img2_5 | 7693.7058
test_results_analysis_values_postcal_standard_3_img2_6 | 6985.6622
test_results_analysis_values_postcal_standard_3_img2_7 | 53949.0556
test_results_analysis_values_postcal_standard_3_img2_8 | 64905.1667
test_results_analysis_values_postcal_standard_3_img2_9 | 55254.7222
test_results_analysis_values_postcal_standard_3_img3_1 | 43.5449
test_results_analysis_values_postcal_standard_3_img3_2 | 8.4658
test_results_analysis_values_postcal_standard_3_img3_3 | 21.7741
test_results_analysis_values_postcal_standard_3_img3_4 | 6959.3053
test_results_analysis_values_postcal_standard_3_img3_5 | 7692.4054
test_results_analysis_values_postcal_standard_3_img3_6 | 6979.7567
test_results_analysis_values_postcal_standard_3_img3_7 | 54030.0556
test_results_analysis_values_postcal_standard_3_img3_8 | 64806.1667
test_results_analysis_values_postcal_standard_3_img3_9 | 55307.7222
test_results_analysis_values_postcal_standard_4_img1_1 | 30396.6914
test_results_analysis_values_postcal_standard_4_img1_2 | 33805.284
test_results_analysis_values_postcal_standard_4_img1_3 | 30135.6806
test_results_analysis_values_postcal_standard_4_img1_4 | 6.2552
test_results_analysis_values_postcal_standard_4_img1_5 | -1.6774
test_results_analysis_values_postcal_standard_4_img1_6 | -0.28067
test_results_analysis_values_postcal_standard_4_img1_7 | 39230.0556
test_results_analysis_values_postcal_standard_4_img1_8 | 39680.1667
test_results_analysis_values_postcal_standard_4_img1_9 | 35066.7222
test_results_analysis_values_postcal_standard_4_img2_1 | 30395.9644
test_results_analysis_values_postcal_standard_4_img2_2 | 33816.328
test_results_analysis_values_postcal_standard_4_img2_3 | 30126.7992
test_results_analysis_values_postcal_standard_4_img2_4 | -5.578
test_results_analysis_values_postcal_standard_4_img2_5 | -2.4662
test_results_analysis_values_postcal_standard_4_img2_6 | 0.48923
test_results_analysis_values_postcal_standard_4_img2_7 | 39136.5556
test_results_analysis_values_postcal_standard_4_img2_8 | 39699.6667
test_results_analysis_values_postcal_standard_4_img2_9 | 35043.2222
test_results_analysis_values_postcal_standard_4_img3_1 | 30392.4504
test_results_analysis_values_postcal_standard_4_img3_2 | 33783.9863
test_results_analysis_values_postcal_standard_4_img3_3 | 30131.3865
test_results_analysis_values_postcal_standard_4_img3_4 | -0.67722
test_results_analysis_values_postcal_standard_4_img3_5 | 4.1436
test_results_analysis_values_postcal_standard_4_img3_6 | -0.20856
test_results_analysis_values_postcal_standard_4_img3_7 | 38936.0556
test_results_analysis_values_postcal_standard_4_img3_8 | 39773.1667
test_results_analysis_values_postcal_standard_4_img3_9 | 34956.7222
test_results_analysis_values_postcal_standard_5_img1_1 | 14129.5469
test_results_analysis_values_postcal_standard_5_img1_2 | 16015.1264
test_results_analysis_values_postcal_standard_5_img1_3 | 13988.94
test_results_analysis_values_postcal_standard_5_img1_4 | -12.9887
test_results_analysis_values_postcal_standard_5_img1_5 | -16.4329
test_results_analysis_values_postcal_standard_5_img1_6 | -37.012
test_results_analysis_values_postcal_standard_5_img1_7 | 18430.5556
test_results_analysis_values_postcal_standard_5_img1_8 | 18993.6667
test_results_analysis_values_postcal_standard_5_img1_9 | 16411.2222
test_results_analysis_values_postcal_standard_5_img2_1 | 14141.8121
test_results_analysis_values_postcal_standard_5_img2_2 | 16026.0289
test_results_analysis_values_postcal_standard_5_img2_3 | 14001.4313
test_results_analysis_values_postcal_standard_5_img2_4 | -10.6326
test_results_analysis_values_postcal_standard_5_img2_5 | -15.038
test_results_analysis_values_postcal_standard_5_img2_6 | -29.9419
test_results_analysis_values_postcal_standard_5_img2_7 | 18412.5556
test_results_analysis_values_postcal_standard_5_img2_8 | 18846.6667
test_results_analysis_values_postcal_standard_5_img2_9 | 16401.2222
test_results_analysis_values_postcal_standard_5_img3_1 | 14126.8329
test_results_analysis_values_postcal_standard_5_img3_2 | 16011.8442
test_results_analysis_values_postcal_standard_5_img3_3 | 13976.46
test_results_analysis_values_postcal_standard_5_img3_4 | -6.678
test_results_analysis_values_postcal_standard_5_img3_5 | -9.6346
test_results_analysis_values_postcal_standard_5_img3_6 | -30.1524
test_results_analysis_values_postcal_standard_5_img3_7 | 18387.5556
test_results_analysis_values_postcal_standard_5_img3_8 | 18708.6667
test_results_analysis_values_postcal_standard_5_img3_9 | 16307.2222
test_results_analysis_values_postcal_standard_6_img1_1 | 6656.216
test_results_analysis_values_postcal_standard_6_img1_2 | 6491.2057
test_results_analysis_values_postcal_standard_6_img1_3 | 6280.4589
test_results_analysis_values_postcal_standard_6_img1_4 | -37.0588
test_results_analysis_values_postcal_standard_6_img1_5 | -11.2977
test_results_analysis_values_postcal_standard_6_img1_6 | -24.1074
test_results_analysis_values_postcal_standard_6_img1_7 | 9429.5556
test_results_analysis_values_postcal_standard_6_img1_8 | 8606.6667
test_results_analysis_values_postcal_standard_6_img1_9 | 8201.2222
test_results_analysis_values_postcal_standard_6_img2_1 | 6646.3781
test_results_analysis_values_postcal_standard_6_img2_2 | 6496.0189
test_results_analysis_values_postcal_standard_6_img2_3 | 6283.1127
test_results_analysis_values_postcal_standard_6_img2_4 | -30.7288
test_results_analysis_values_postcal_standard_6_img2_5 | -8.9411
test_results_analysis_values_postcal_standard_6_img2_6 | -16.054
test_results_analysis_values_postcal_standard_6_img2_7 | 9578.5556
test_results_analysis_values_postcal_standard_6_img2_8 | 8566.6667
test_results_analysis_values_postcal_standard_6_img2_9 | 8146.2222
test_results_analysis_values_postcal_standard_6_img3_1 | 6646.9031
test_results_analysis_values_postcal_standard_6_img3_2 | 6501.9354
test_results_analysis_values_postcal_standard_6_img3_3 | 6269.4009
test_results_analysis_values_postcal_standard_6_img3_4 | -26.6793
test_results_analysis_values_postcal_standard_6_img3_5 | -6.4966
test_results_analysis_values_postcal_standard_6_img3_6 | -12.8854
test_results_analysis_values_postcal_standard_6_img3_7 | 9470.0556
test_results_analysis_values_postcal_standard_6_img3_8 | 8632.1667
test_results_analysis_values_postcal_standard_6_img3_9 | 8355.7222
test_results_analysis_values_postcal_standard_7_img1_1 | 4.2951
test_results_analysis_values_postcal_standard_7_img1_2 | -10.1115
test_results_analysis_values_postcal_standard_7_img1_3 | 7.9328
test_results_analysis_values_postcal_standard_7_img1_4 | -16.7071
test_results_analysis_values_postcal_standard_7_img1_5 | -8.2873
test_results_analysis_values_postcal_standard_7_img1_6 | -8.8929
test_results_analysis_values_postcal_standard_7_img1_7 | 160.0556
test_results_analysis_values_postcal_standard_7_img1_8 | 160.1667
test_results_analysis_values_postcal_standard_7_img1_9 | 253.7222
test_results_analysis_values_postcal_standard_7_img2_1 | 8.8216
test_results_analysis_values_postcal_standard_7_img2_2 | -3.4829
test_results_analysis_values_postcal_standard_7_img2_3 | 12.1115
test_results_analysis_values_postcal_standard_7_img2_4 | -14.8484
test_results_analysis_values_postcal_standard_7_img2_5 | -7.2761
test_results_analysis_values_postcal_standard_7_img2_6 | -4.3011
test_results_analysis_values_postcal_standard_7_img2_7 | 35.0556
test_results_analysis_values_postcal_standard_7_img2_8 | 202.1667
test_results_analysis_values_postcal_standard_7_img2_9 | 290.7222
test_results_analysis_values_postcal_standard_7_img3_1 | 10.0185
test_results_analysis_values_postcal_standard_7_img3_2 | -6.4329
test_results_analysis_values_postcal_standard_7_img3_3 | 9.761
test_results_analysis_values_postcal_standard_7_img3_4 | -17.2251
test_results_analysis_values_postcal_standard_7_img3_5 | -4.7963
test_results_analysis_values_postcal_standard_7_img3_6 | -6.7195
test_results_analysis_values_postcal_standard_7_img3_7 | 187.0556
test_results_analysis_values_postcal_standard_7_img3_8 | 324.1667
test_results_analysis_values_postcal_standard_7_img3_9 | 297.7222
test_results_analysis_values_postcal_standard_8_img1_1 | 44.6902
test_results_analysis_values_postcal_standard_8_img1_2 | 21.0249
test_results_analysis_values_postcal_standard_8_img1_3 | 3.0759
test_results_analysis_values_postcal_standard_8_img1_4 | 22.9856
test_results_analysis_values_postcal_standard_8_img1_5 | 8.0135
test_results_analysis_values_postcal_standard_8_img1_6 | 5.9713
test_results_analysis_values_postcal_standard_8_img1_7 | 680.0556
test_results_analysis_values_postcal_standard_8_img1_8 | 431.1667
test_results_analysis_values_postcal_standard_8_img1_9 | 410.7222
test_results_analysis_values_postcal_standard_8_img2_1 | -32.0434
test_results_analysis_values_postcal_standard_8_img2_2 | -61.7596
test_results_analysis_values_postcal_standard_8_img2_3 | -79.998
test_results_analysis_values_postcal_standard_8_img2_4 | 5.9546
test_results_analysis_values_postcal_standard_8_img2_5 | -9.1977
test_results_analysis_values_postcal_standard_8_img2_6 | -5.9472
test_results_analysis_values_postcal_standard_8_img2_7 | -686.4444
test_results_analysis_values_postcal_standard_8_img2_8 | -974.3333
test_results_analysis_values_postcal_standard_8_img2_9 | -972.7778
test_results_analysis_values_postcal_standard_8_img3_1 | -23.1424
test_results_analysis_values_postcal_standard_8_img3_2 | -55.5204
test_results_analysis_values_postcal_standard_8_img3_3 | -75.2525
test_results_analysis_values_postcal_standard_8_img3_4 | 13.2131
test_results_analysis_values_postcal_standard_8_img3_5 | -7.4665
test_results_analysis_values_postcal_standard_8_img3_6 | -7.4589
test_results_analysis_values_postcal_standard_8_img3_7 | -646.9444
test_results_analysis_values_postcal_standard_8_img3_8 | -802.8333
test_results_analysis_values_postcal_standard_8_img3_9 | -848.2778
test_results_analysis_values_postcal_standard_9_img1_1 | -44.0365
test_results_analysis_values_postcal_standard_9_img1_2 | -4.109
test_results_analysis_values_postcal_standard_9_img1_3 | 1.1314
test_results_analysis_values_postcal_standard_9_img1_4 | 24.1128
test_results_analysis_values_postcal_standard_9_img1_5 | 30.6577
test_results_analysis_values_postcal_standard_9_img1_6 | 30.7965
test_results_analysis_values_postcal_standard_9_img1_7 | 588.0556
test_results_analysis_values_postcal_standard_9_img1_8 | 693.1667
test_results_analysis_values_postcal_standard_9_img1_9 | 603.7222
test_results_analysis_values_postcal_standard_9_img2_1 | 57.3167
test_results_analysis_values_postcal_standard_9_img2_2 | 101.7249
test_results_analysis_values_postcal_standard_9_img2_3 | 96.3312
test_results_analysis_values_postcal_standard_9_img2_4 | -13.3092
test_results_analysis_values_postcal_standard_9_img2_5 | -4.5507
test_results_analysis_values_postcal_standard_9_img2_6 | -6.4675
test_results_analysis_values_postcal_standard_9_img2_7 | -836.4444
test_results_analysis_values_postcal_standard_9_img2_8 | -760.3333
test_results_analysis_values_postcal_standard_9_img2_9 | -721.7778
test_results_analysis_values_postcal_standard_9_img3_1 | -25.9199
test_results_analysis_values_postcal_standard_9_img3_2 | 18.6665
test_results_analysis_values_postcal_standard_9_img3_3 | 24.9068
test_results_analysis_values_postcal_standard_9_img3_4 | -4.1762
test_results_analysis_values_postcal_standard_9_img3_5 | 2.9034
test_results_analysis_values_postcal_standard_9_img3_6 | 3.0192
test_results_analysis_values_postcal_standard_9_img3_7 | 519.5556
test_results_analysis_values_postcal_standard_9_img3_8 | 726.6667
test_results_analysis_values_postcal_standard_9_img3_9 | 686.2222
test_results_cal_values_dark_well_1 | 0
test_results_cal_values_dark_well_2 | 0
test_results_cal_values_dark_well_3 | 0
test_results_cal_values_background_1 | 14350.1769
test_results_cal_values_background_2 | 5451.5648
test_results_cal_values_background_3 | 13215.0808
test_results_cal_values_background_4 | 10454.494
test_results_cal_values_background_5 | 1133.6985
test_results_cal_values_background_6 | 9467.051
test_results_cal_values_background_7 | 0
test_results_cal_values_background_8 | 0
test_results_cal_values_background_9 | 0
test_results_cal_values_normalizers_r_norm_1 | 16.9801
test_results_cal_values_normalizers_r_norm_2 | 15.69
test_results_cal_values_normalizers_r_norm_3 | 16.459
test_results_cal_values_normalizers_g_norm_1 | 23.7047
test_results_cal_values_normalizers_g_norm_2 | 28.2839
test_results_cal_values_normalizers_g_norm_3 | 25.2398
test_results_cal_values_crosstalks_xrgs_1 | 0.188
test_results_cal_values_crosstalks_xrgs_2 | 0.17178
test_results_cal_values_crosstalks_xrgs_3 | 0.17155
test_results_cal_values_crosstalks_xgrs_1 | 0.17655
test_results_cal_values_crosstalks_xgrs_2 | 0.17758
test_results_cal_values_crosstalks_xgrs_3 | 0.19742
test_results_ref_values_red_1 | 30432.9104
test_results_ref_values_red_2 | 33696.0392
test_results_ref_values_red_3 | 30077.4416
test_results_ref_values_green_1 | 31552.1055
test_results_ref_values_green_2 | 32675.1877
test_results_ref_values_green_3 | 30245.7547
test_results_spec_values_spec_results_well_1_lin_green | 0.99802
test_results_spec_values_spec_results_well_1_lin_red | 0.99959
test_results_spec_values_spec_results_well_1_standard_2_valg | 0.0010395
test_results_spec_values_spec_results_well_1_standard_3_valg | 0.0054632
test_results_spec_values_spec_results_well_1_standard_5_valr | -0.00071464
test_results_spec_values_spec_results_well_1_standard_6_valr | -0.0047353
test_results_spec_values_spec_results_well_2_lin_green | 0.99952
test_results_spec_values_spec_results_well_2_lin_red | 0.99971
test_results_spec_values_spec_results_well_2_standard_2_valg | -0.0025123
test_results_spec_values_spec_results_well_2_standard_3_valg | 0.00081067
test_results_spec_values_spec_results_well_2_standard_5_valr | -0.00085542
test_results_spec_values_spec_results_well_2_standard_6_valr | -0.0013718
test_results_spec_values_spec_results_well_3_lin_green | 0.99978
test_results_spec_values_spec_results_well_3_lin_red | 0.99988
test_results_spec_values_spec_results_well_3_standard_2_valg | -0.00045004
test_results_spec_values_spec_results_well_3_standard_3_valg | 0.0028933
test_results_spec_values_spec_results_well_3_standard_5_valr | -0.0023139
test_results_spec_values_spec_results_well_3_standard_6_valr | -0.0028167
test_results_spec_values_data_normalizers_r_norm_1 | 16.9801
test_results_spec_values_data_normalizers_r_norm_2 | 15.69
test_results_spec_values_data_normalizers_r_norm_3 | 16.459
test_results_spec_values_data_normalizers_g_norm_1 | 23.7047
test_results_spec_values_data_normalizers_g_norm_2 | 28.2839
test_results_spec_values_data_normalizers_g_norm_3 | 25.2398
test_results_spec_values_data_crosstalks_xrgs_1 | 0.188
test_results_spec_values_data_crosstalks_xrgs_2 | 0.17178
test_results_spec_values_data_crosstalks_xrgs_3 | 0.17155
test_results_spec_values_data_crosstalks_xgrs_1 | 0.17655
test_results_spec_values_data_crosstalks_xgrs_2 | 0.17758
test_results_spec_values_data_crosstalks_xgrs_3 | 0.19742
test_results_version_version_number | 2
test_results_version_info | Initial R3 TEMP HACK

## optical_targets
batchcode | 202007240108
ref_unit | B19460006
timestamp | 20210304-2025
ref_cal_vals_1_1 | -111.2651
ref_cal_vals_1_2 | -156.1821
ref_cal_vals_1_3 | -148.6273
ref_cal_vals_1_4 | 31552.1055
ref_cal_vals_1_5 | 32675.1877
ref_cal_vals_1_6 | 30245.7547
ref_cal_vals_2_1 | -85.5476
ref_cal_vals_2_2 | -212.5217
ref_cal_vals_2_3 | -257.7427
ref_cal_vals_2_4 | 16213.0861
ref_cal_vals_2_5 | 16322.6945
ref_cal_vals_2_6 | 15356.9713
ref_cal_vals_3_1 | -111.2139
ref_cal_vals_3_2 | -229.4637
ref_cal_vals_3_3 | -300.6165
ref_cal_vals_3_4 | 6727.4892
ref_cal_vals_3_5 | 7739.0747
ref_cal_vals_3_6 | 7202.2284
ref_cal_vals_4_1 | 30432.9104
ref_cal_vals_4_2 | 33696.0392
ref_cal_vals_4_3 | 30077.4416
ref_cal_vals_4_4 | -237.5394
ref_cal_vals_4_5 | -211.744
ref_cal_vals_4_6 | -290.8598
ref_cal_vals_5_1 | 14191.0297
ref_cal_vals_5_2 | 16213.1043
ref_cal_vals_5_3 | 13806.2851
ref_cal_vals_5_4 | -175.8102
ref_cal_vals_5_5 | -158.6836
ref_cal_vals_5_6 | -198.7036
ref_cal_vals_6_1 | 6598.9444
ref_cal_vals_6_2 | 6094.7451
ref_cal_vals_6_3 | 6107.8198
ref_cal_vals_6_4 | -169.1524
ref_cal_vals_6_5 | -213.638
ref_cal_vals_6_6 | -91.553
ref_cal_vals_7_1 | -135.2569
ref_cal_vals_7_2 | -235.8671
ref_cal_vals_7_3 | -312.4438
ref_cal_vals_7_4 | -117.8029
ref_cal_vals_7_5 | -188.6153
ref_cal_vals_7_6 | 9.4243
ref_cal_vals_8_1 | -116.4056
ref_cal_vals_8_2 | -214.8377
ref_cal_vals_8_3 | -332.6426
ref_cal_vals_8_4 | -110.7137
ref_cal_vals_8_5 | -183.4026
ref_cal_vals_8_6 | -11.4689
ref_cal_vals_9_1 | -105.8805
ref_cal_vals_9_2 | -148.319
ref_cal_vals_9_3 | -245.7728
ref_cal_vals_9_4 | -122.1307
ref_cal_vals_9_5 | -174.0747
ref_cal_vals_9_6 | -4.7726
green_targets_1 | 31552.1055
green_targets_2 | 32675.1877
green_targets_3 | 30245.7547
red_targets_1 | 30432.9104
red_targets_2 | 33696.0392
red_targets_3 | 30077.4416